Data/hora: 09/06/2024 03:57:55

Provedor de dados: 	5	
País: 	France	
Título: 	Recommendations for obtaining unbiased chlorophyll estimates from in situ chlorophyll fluorometers: A global analysis of WET Labs ECO sensors	
Autores: 	Roesler, Collin; Uitz, Julia; Claustre, Herve; Boss, Emmanuel; Xing, Xiaogang; Organelli, Emanuele; Briggs, Nathan; Bricaud, Annick; Schmechtig, Catherine; Poteau, Antoine; D'Ortenzio, Fabrizio; Ras, Josephine; Drapeau, Susan; Haentjens, Nils; Barbieux, Marie.	
Data: 	2017-06	
Ano: 	2017	
Resumo: 	Chlorophyll fluorometers provide the largest in situ global data set for estimating phytoplankton biomass because of their ease of use, size, power consumption, and relatively low price. While in situ chlorophyll a (Chl) fluorescence is proxy for Chl a concentration, and hence phytoplankton biomass, there exist large natural variations in the relationship between in situ fluorescence and extracted Chl a concentration. Despite this large natural variability, we present here a global validation data set for the WET Labs Environmental Characterization Optics (ECO) series chlorophyll fluorometers that suggests a factor of 2 overestimation in the factory calibrated Chl a estimates for this specific manufacturer and series of sensors. We base these results on paired High Pressure Liquid Chromatography (HPLC) and in situ fluorescence match ups for which non-photochemically quenched fluorescence observations were removed. Additionally, we examined matchups between the factory-calibrated in situ fluorescence and estimates of chlorophyll concentration determined from in situ radiometry, absorption line height, NASA's standard ocean color algorithm as well as laboratory calibrations with phytoplankton monocultures spanning diverse species that support the factor of 2 bias. We therefore recommend the factor of 2 global bias correction be applied for the WET Labs ECO sensors, at the user level, to improve the global accuracy of chlorophyll concentration estimates and products derived from them. We recommend that other fluorometer makes and models should likewise undergo global analyses to identify potential bias in factory calibration.	
Tipo: 	Text	
Idioma: 	Inglês	
Identificador: 	https://archimer.ifremer.fr/doc/00384/49552/50046.pdf

DOI:10.1002/lom3.10185

https://archimer.ifremer.fr/doc/00384/49552/	
Editor: 	Wiley	
Formato: 	application/pdf	
Fonte: 	Limnology And Oceanography-methods (1541-5856) (Wiley), 2017-06 , Vol. 15 , N. 6 , P. 572-585	
Fonte: 	2017 The Authors Limnology and Oceanography: Methods published by Wiley Periodicals, Inc. on behalf of Association for the Sciences of Limnology and Oceanography This is an open access article under the terms of the Creative Commons Attribution License, which permits use, distribution and reproduction in any medium, provided the original work is properly cited.

info:eu-repo/semantics/openAccess

restricted use