Data/hora: 13/06/2024 04:10:14

Provedor de dados: | 56
País: | Brazil
Título: | Performance assessment of the single photon emission microscope: high spatial resolution SPECT imaging of small animal organs
Autores: | Mejia,J.; Reis,M.A.; Miranda,A.C.C.; Batista,I.R.; Barboza,M.R.F.; Shih,M.C.; Fu,G.; Chen,C.T.; Meng,L.J.; Bressan,R.A.; Amaro Jr,E.
Data: | 2013-10-01
Ano: | 2013
Palavras-chave: | Tomography; Preclinical imaging; Molecular imaging; Mice.
Resumo: | The single photon emission microscope (SPEM) is an instrument developed to obtain high spatial resolution single photon emission computed tomography (SPECT) images of small structures inside the mouse brain. SPEM consists of two independent imaging devices, which combine a multipinhole collimator, a high-resolution, thallium-doped cesium iodide [CsI(Tl)] columnar scintillator, a demagnifying/intensifier tube, and an electron-multiplying charge-coupling device (CCD). Collimators have 300- and 450-µm diameter pinholes on tungsten slabs, in hexagonal arrays of 19 and 7 holes. Projection data are acquired in a photon-counting strategy, where CCD frames are stored at 50 frames per second, with a radius of rotation of 35 mm and magnification factor of one. The image reconstruction software tool is based on the maximum likelihood algorithm. Our aim was to evaluate the spatial resolution and sensitivity attainable with the seven-pinhole imaging device, together with the linearity for quantification on the tomographic images, and to test the instrument in obtaining tomographic images of different mouse organs. A spatial resolution better than 500 µm and a sensitivity of 21.6 counts·s-1·MBq-1 were reached, as well as a correlation coefficient between activity and intensity better than 0.99, when imaging 99mTc sources. Images of the thyroid, heart, lungs, and bones of mice were registered using 99mTc-labeled radiopharmaceuticals in times appropriate for routine preclinical experimentation of &lt;1 h per projection data set. Detailed experimental protocols and images of the aforementioned organs are shown. We plan to extend the instrument's field of view to fix larger animals and to combine data from both detectors to reduce the acquisition time or applied activity.
Tipo: | Info:eu-repo/semantics/article
Idioma: | Inglês
Identificador: | http://www.scielo.br/scielo.php?script=sci_arttext&pid=S0100-879X2013001100936
Editor: | Associação Brasileira de Divulgação Científica
Relação: | 10.1590/1414-431X20132764
Formato: | text/html
Fonte: | Brazilian Journal of Medical and Biological Research  v.46 n.11 2013
Fonte: | info:eu-repo/semantics/openAccess