Data/hora: 05/06/2024 02:31:58

Provedor de dados: 	56	
País: 	Brazil	
Título: 	Electroretinography in dogs using a fiber electrode prototype	
Autores: 	Pereira,A.L.; Montiani-Ferreira,F.; Santos,V.R.; Salomão,S.R.; Souza,C.; Berezovsky,A.	
Data: 	2013-03-01	
Ano: 	2013	
Palavras-chave: 	Electroretinography; Electrode; Retina; Canine; Prototype.	
Resumo: 	We compared two electroretinography (ERG) electrodes in dogs using ERG standards of the International Society for Clinical Electrophysiology of Vision (ISCEV). Ten healthy Yorkshire terrier dogs (mean age, 2.80 ± 1.42 years; 6 females) weighing 5.20 ± 1.56 kg were evaluated using an ERG system for veterinary use. Dark- and light-adapted ERG responses were recorded using an ERG-Jet electrode and a fiber electrode prototype. The examinations were performed during 2 visits, 3 weeks apart. Both electrodes (ERG-Jet or fiber prototype) were used on each animal and the first eye to be recorded (OD × OS) was selected randomly. Three weeks later the examination was repeated on the same animal switching the type of electrode to be used that day and the first eye to be examined. The magnitude and waveform quality obtained with the two electrode types were similar for all ERG responses. ERG amplitudes and implicit times obtained from dogs using the fiber electrode prototype were comparable to those obtained with the ERG-Jet electrode for rod, maximal rod-cone summed, cone, and 30-Hz flicker responses. The fiber electrode prototype is a low-cost device, available as an alternative instrument for clinical veterinary ERG recording for retinal function assessment.	
Tipo: 	Info:eu-repo/semantics/article	
Idioma: 	Inglês	
Identificador: 	http://www.scielo.br/scielo.php?script=sci_arttext&pid=S0100-879X2013000300257	
Editor: 	Associação Brasileira de Divulgação Científica	
Relação: 	10.1590/1414-431X20122769	
Formato: 	text/html	
Fonte: 	Brazilian Journal of Medical and Biological Research  v.46 n.3 2013	
Fonte: 	info:eu-repo/semantics/openAccess